## SL0NormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
Single Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## AIINormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
AI Infant Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.001 | 0.002 | 0.003 | 0.004 | 0.005 | 0.006 | 0.007 | 0.008 | 0.009 | 0.01
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.001 to cmH2O 0.01
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.001: Compliance Setting
| Volume | 0.001: (L/cmH2O)
| (mL) | 0.001 | 0.002 | 0.003 | 0.004 | 0.005 | 0.006 | 0.007 | 0.008 | 0.009 | 0.01
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 10 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 20 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-08
| 30 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -1.1907e-05
| 40 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.0019629
| 50 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 0.89895
| 60 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 70 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 80 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-08 | -1.1907e-05 | 0.0019629 | 0.89895 | 0
| 90 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-09 | 2.6416e-06 | 0.00031764 | 0.011914 | #N/A
| 100 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 0.31615 | #N/A | #N/A | #N/A
| 110 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 120 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 286998.3009 | 1.6991 | #N/A | #N/A | #N/A
| 130 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 140 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 150 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 160 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 170 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 180 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 190 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-07
| 200 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.00021852
[12 empty] | CC[2] | 0.01736
[12 empty] | CC[3] | 1.5614
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.001: Compliance Setting
| Volume | (L/cmH2O) 0.001: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-06
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.0011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.058651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 2.0363
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 2.588e-05
[12 empty] | CC[1] | -0.0030293
[12 empty] | CC[2] | 0.1185
[12 empty] | CC[3] | 2.5371
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 8.1187e-05
[12 empty] | CC[1] | -0.0074691
[12 empty] | CC[2] | 0.23011
[12 empty] | CC[3] | 2.7441
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.0001967
[12 empty] | CC[1] | -0.015047
[12 empty] | CC[2] | 0.38352
[12 empty] | CC[3] | 2.9139
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.0003779
[12 empty] | CC[1] | -0.02522
[12 empty] | CC[2] | 0.55826
[12 empty] | CC[3] | 3.0974
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.00038442
[12 empty] | CC[1] | -0.028391
[12 empty] | CC[2] | 0.64846
[12 empty] | CC[3] | 3.6
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.0012231
[12 empty] | CC[1] | -0.064082
[12 empty] | CC[2] | 1.1136
[12 empty] | CC[3] | 2.9037
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0020917
[12 empty] | CC[1] | -0.098322
[12 empty] | CC[2] | 1.5276
[12 empty] | CC[3] | 2.5019
[12 empty] | R2 | 0.99993
[13 empty] | 0.0020917 | -0.098322 | 1.5276 | 2.5019 | 0
[13 empty] | 0.00027082 | 0.010151 | 0.11916 | 0.43518 | #N/A
[13 empty] | 0.99993 | 1.0728 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 286980.4352 | 19.5648 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## AIANormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
AI Adult Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## DALNormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
DA Left Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## DARNormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
DA Right Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Validation Data
Validation Data
Adult Lung: Measured Lung Volume
| Injected | Compliance Setting
| Volume | (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
0
100
200
300
400
500
600
700
800
900
1000
1100
1200
1300
1400
1500
1600
1700
1800
1900
2000