## SL0NormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
Single Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1 |  |  |  | Residuals Analysis
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients |  |  | MIITTLCalibration | MathCAD
| 100 | 12.051 | 6.96 | 4.906 | 3.83 | 3.235 | 2.764 | 2.464 | 2.204 | 1.985 | 1.838
| 200 | 21.6515 | 11.9922 | 8.1745 | 6.2775 | 5.2507 | 4.4005 | 3.868 | 3.433 | 3.0875 | 2.864 | CC[0] | 4.4295e-07 |  | 4.9217e-07 | 7.8415
| 300 | 31.4808 | 17.0324 | 11.443 | 8.725 | 7.2185 | 6.037 | 5.272 | 4.662 | 4.19 | 3.8226 | CC[1] | -0.00020015 |  | -0.00021661 | 0.067206
| 400 | 41.5076 | 21.9898 | 14.7115 | 11.1725 | 9.18 | 7.6735 | 6.676 | 5.891 | 5.2925 | 4.7992 | CC[2] | 0.029269 |  | 0.030853 | -0.00063061
| 500 | 51.625 | 26.96 | 17.98 | 13.62 | 11.095 | 9.31 | 8.08 | 7.12 | 6.395 | 5.71 | CC[3] | 8.6754 |  | 8.907 | 1.9067e-06
| 600 | 61.3013 | 31.6884 | 21.03 | 15.881 | 12.8929 | 10.794 | 9.349 | 8.222 | 7.379 | 6.613 | R2 | 0.99999 | 96.57 | 995.6031 | 981.8982
| 700 | 70.8586 | 36.3831 | 24.08 | 18.142 | 14.6693 | 12.278 | 10.618 | 9.324 | 8.363 | 7.4557
| 800 | 80.2256 | 41.0367 | 27.13 | 20.403 | 16.4425 | 13.762 | 11.887 | 10.426 | 9.347 | 8.2896 |  | 4.4295e-07 | -0.00020015 | 0.029269 | 8.6754 | 0
| 900 | 89.6482 | 45.9019 | 30.18 | 22.664 | 18.277 | 15.246 | 13.156 | 11.528 | 10.331 | 9.1774 |  | 6.6845e-08 | 2.5693e-05 | 0.0030915 | 0.11592 | #N/A
| 1000 | 99.15 | 50.57 | 33.23 | 24.925 | 20.11 | 16.73 | 14.425 | 12.63 | 11.315 | 10.045 |  | 0.99999 | 3.0669 | #N/A | #N/A | #N/A
| 1100 | 109.32 | 55.304 | 36.31 | 27.2002 | 21.926 | 18.241 | 15.716 | 13.761 | 12.317 | 10.935 |  | 762810.0488 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.49 | 60.038 | 39.39 | 29.4754 | 23.742 | 19.752 | 17.007 | 14.892 | 13.319 | 11.825 |  | 28699840.0987 | 159.9013 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.772 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 6.7671e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.001763
[12 empty] | CC[2] | 0.16125
[12 empty] | CC[3] | 15.2801
[12 empty] | R2 | 0.99999
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 12.051 | 6.96 | 4.906 | 3.83 | 3.235 | 2.764 | 2.464 | 2.204 | 1.985 | 1.838
| 200 | 21.6515 | 11.9922 | 8.1745 | 6.2775 | 5.2507 | 4.4005 | 3.868 | 3.433 | 3.0875 | 2.864
| 300 | 31.4808 | 17.0324 | 11.443 | 8.725 | 7.2185 | 6.037 | 5.272 | 4.662 | 4.19 | 3.8226
| 400 | 41.5076 | 21.9898 | 14.7115 | 11.1725 | 9.18 | 7.6735 | 6.676 | 5.891 | 5.2925 | 4.7992
| 500 | 51.625 | 26.96 | 17.98 | 13.62 | 11.095 | 9.31 | 8.08 | 7.12 | 6.395 | 5.71
| 600 | 61.3013 | 31.6884 | 21.03 | 15.881 | 12.8929 | 10.794 | 9.349 | 8.222 | 7.379 | 6.613 | Compliance Coefficients
| 700 | 70.8586 | 36.3831 | 24.08 | 18.142 | 14.6693 | 12.278 | 10.618 | 9.324 | 8.363 | 7.4557
| 800 | 80.2256 | 41.0367 | 27.13 | 20.403 | 16.4425 | 13.762 | 11.887 | 10.426 | 9.347 | 8.2896 | CC[0] | 7.0638e-05
| 900 | 89.6482 | 45.9019 | 30.18 | 22.664 | 18.277 | 15.246 | 13.156 | 11.528 | 10.331 | 9.1774 | CC[1] | -0.011021
| 1000 | 99.15 | 50.57 | 33.23 | 24.925 | 20.11 | 16.73 | 14.425 | 12.63 | 11.315 | 10.045 | CC[2] | 0.57509
| 1100 | 109.32 | 55.304 | 36.31 | 27.2002 | 21.926 | 18.241 | 15.716 | 13.761 | 12.317 | 10.935 | CC[3] | 20.5898
| 1200 | 119.49 | 60.038 | 39.39 | 29.4754 | 23.742 | 19.752 | 17.007 | 14.892 | 13.319 | 11.825 | R2 | 0.99999
| 1300 | 129.3083 | 64.772 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.00024096
[12 empty] | CC[1] | -0.028579
[12 empty] | CC[2] | 1.1305
[12 empty] | CC[3] | 25.9923
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.00064148
[12 empty] | CC[1] | -0.061225
[12 empty] | CC[2] | 1.9756
[12 empty] | CC[3] | 29.6939
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.0012058
[12 empty] | CC[1] | -0.10032
[12 empty] | CC[2] | 2.8187
[12 empty] | CC[3] | 35.1455
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.002031
[12 empty] | CC[1] | -0.15372
[12 empty] | CC[2] | 3.8884
[12 empty] | CC[3] | 39.2947
[12 empty] | R2 | 0.99998
Compliance Coefficients
[12 empty] | CC[0] | 0.00041575
[12 empty] | CC[1] | -0.10787
[12 empty] | CC[2] | 3.7046
[12 empty] | CC[3] | 48.7833
[12 empty] | R2 | 0.99994
Compliance Coefficients
[12 empty] | CC[0] | 0.0037776
[12 empty] | CC[1] | -0.2648
[12 empty] | CC[2] | 5.9918
[12 empty] | CC[3] | 49.4358
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.011211
[12 empty] | CC[1] | -0.58058
[12 empty] | CC[2] | 9.9759
[12 empty] | CC[3] | 46.7806
[12 empty] | R2 | 0.99998
[13 empty] | 0.011211 | -0.58058 | 9.9759 | 46.7806 | 0
[13 empty] | 0.0013357 | 0.050038 | 0.5877 | 2.1546 | #N/A
[13 empty] | 0.99998 | 5.3568 | #N/A | #N/A | #N/A
[13 empty] | 250041.2211 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28699512.1887 | 487.8113 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## AIINormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
AI Infant Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.001 | 0.002 | 0.003 | 0.004 | 0.005 | 0.006 | 0.007 | 0.008 | 0.009 | 0.01
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.001 to cmH2O 0.01
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.001: Compliance Setting
| Volume | 0.001: (L/cmH2O)
| (mL) | 0.001 | 0.002 | 0.003 | 0.004 | 0.005 | 0.006 | 0.007 | 0.008 | 0.009 | 0.01
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 10 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 20 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-08
| 30 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -1.1907e-05
| 40 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.0019629
| 50 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 0.89895
| 60 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 70 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 80 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-08 | -1.1907e-05 | 0.0019629 | 0.89895 | 0
| 90 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-09 | 2.6416e-06 | 0.00031764 | 0.011914 | #N/A
| 100 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 0.31615 | #N/A | #N/A | #N/A
| 110 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 120 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 286998.3009 | 1.6991 | #N/A | #N/A | #N/A
| 130 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 140 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 150 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 160 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 170 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 180 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 190 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-07
| 200 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.00021852
[12 empty] | CC[2] | 0.01736
[12 empty] | CC[3] | 1.5614
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.001: Compliance Setting
| Volume | (L/cmH2O) 0.001: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-06
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.0011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.058651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 2.0363
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 2.588e-05
[12 empty] | CC[1] | -0.0030293
[12 empty] | CC[2] | 0.1185
[12 empty] | CC[3] | 2.5371
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 8.1187e-05
[12 empty] | CC[1] | -0.0074691
[12 empty] | CC[2] | 0.23011
[12 empty] | CC[3] | 2.7441
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.0001967
[12 empty] | CC[1] | -0.015047
[12 empty] | CC[2] | 0.38352
[12 empty] | CC[3] | 2.9139
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.0003779
[12 empty] | CC[1] | -0.02522
[12 empty] | CC[2] | 0.55826
[12 empty] | CC[3] | 3.0974
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.00038442
[12 empty] | CC[1] | -0.028391
[12 empty] | CC[2] | 0.64846
[12 empty] | CC[3] | 3.6
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.0012231
[12 empty] | CC[1] | -0.064082
[12 empty] | CC[2] | 1.1136
[12 empty] | CC[3] | 2.9037
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0020917
[12 empty] | CC[1] | -0.098322
[12 empty] | CC[2] | 1.5276
[12 empty] | CC[3] | 2.5019
[12 empty] | R2 | 0.99993
[13 empty] | 0.0020917 | -0.098322 | 1.5276 | 2.5019 | 0
[13 empty] | 0.00027082 | 0.010151 | 0.11916 | 0.43518 | #N/A
[13 empty] | 0.99993 | 1.0728 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 286980.4352 | 19.5648 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## AIANormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
AI Adult Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 12.728 | 7.302 | 5.108 | 4.018 | 3.334 | 2.826 | 2.458 | 2.1272 | 1.982 | 1.7272
| 200 | 22.1448 | 12.2549 | 8.386 | 6.4885 | 5.3665 | 4.487 | 3.8835 | 3.3744 | 3.0915 | 2.7836 | CC[0] | 5.7414e-07
| 300 | 31.8164 | 17.2168 | 11.664 | 8.959 | 7.3436 | 6.148 | 5.309 | 4.6216 | 4.201 | 3.7595 | CC[1] | -0.00025405
| 400 | 41.7078 | 22.0862 | 14.942 | 11.4295 | 9.3136 | 7.809 | 6.7345 | 5.8688 | 5.3105 | 4.757 | CC[2] | 0.035914
| 500 | 51.7 | 26.97 | 18.22 | 13.9 | 11.23 | 9.47 | 8.16 | 7.116 | 6.42 | 5.676 | CC[3] | 8.4017
| 600 | 61.4368 | 31.6571 | 21.258 | 16.164 | 13.0362 | 10.976 | 9.442 | 8.2256 | 7.4062 | 6.5886 | R2 | 0.99999
| 700 | 71.0397 | 36.3068 | 24.296 | 18.428 | 14.8182 | 12.482 | 10.724 | 9.3352 | 8.3924 | 7.4336
| 800 | 80.4284 | 40.911 | 27.334 | 20.692 | 16.5966 | 13.988 | 12.006 | 10.4448 | 9.3786 | 8.2686 |  | 5.7414e-07 | -0.00025405 | 0.035914 | 8.4017 | 0
| 900 | 89.8796 | 45.7497 | 30.372 | 22.956 | 18.4442 | 15.494 | 13.288 | 11.5544 | 10.3648 | 9.1641 |  | 8.0863e-08 | 3.1228e-05 | 0.0037743 | 0.14211 | #N/A
| 1000 | 99.42 | 50.37 | 33.41 | 25.22 | 20.29 | 17 | 14.57 | 12.664 | 11.351 | 10.037 |  | 0.99999 | 3.8002 | #N/A | #N/A | #N/A
| 1100 | 109.536 | 55.03 | 36.488 | 27.528 | 22.118 | 18.54 | 15.89 | 13.8052 | 12.3628 | 10.9596 |  | 496815.0137 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.652 | 59.69 | 39.566 | 29.836 | 23.946 | 20.08 | 17.21 | 14.9464 | 13.3746 | 11.8822 |  | 28699754.4882 | 245.5118 | #N/A | #N/A | #N/A
| 1300 | 129.768 | 64.35 | 42.644 | 32.144 | 25.774 | 21.62 | 18.53 | 16.0876 | 14.3864 | 12.8048
| 1400 | 139.884 | 69.01 | 45.722 | 34.452 | 27.602 | 23.16 | 19.85 | 17.2288 | 15.3982 | 13.7274
| 1500 | 150 | 73.67 | 48.8 | 36.76 | 29.43 | 24.7 | 21.17 | 18.37 | 16.41 | 14.65
| 1600 | 160 | 78.382 | 51.926 | 39.112 | 31.35 | 26.318 | 22.536 | 19.598 | 17.496 | 15.758
| 1700 | 170 | 83.094 | 55.052 | 41.464 | 33.27 | 27.936 | 23.902 | 20.826 | 18.582 | 16.866 | Compliance Coefficients
| 1800 | 180 | 87.806 | 58.178 | 43.816 | 35.19 | 29.554 | 25.268 | 22.054 | 19.668 | 17.974
| 1900 | 190 | 92.518 | 61.304 | 46.168 | 37.11 | 31.172 | 26.634 | 23.282 | 20.754 | 19.082 | CC[0] | 1.0921e-05
| 2000 | 200 | 97.23 | 64.43 | 48.52 | 39.03 | 32.79 | 28 | 24.51 | 21.84 | 20.19 | CC[1] | -0.0025893
[12 empty] | CC[2] | 0.21271
[12 empty] | CC[3] | 14.3634
[12 empty] | R2 | 0.99999
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 12.728 | 7.302 | 5.108 | 4.018 | 3.334 | 2.826 | 2.458 | 2.1272 | 1.982 | 1.7272
| 200 | 22.1448 | 12.2549 | 8.386 | 6.4885 | 5.3665 | 4.487 | 3.8835 | 3.3744 | 3.0915 | 2.7836
| 300 | 31.8164 | 17.2168 | 11.664 | 8.959 | 7.3436 | 6.148 | 5.309 | 4.6216 | 4.201 | 3.7595
| 400 | 41.7078 | 22.0862 | 14.942 | 11.4295 | 9.3136 | 7.809 | 6.7345 | 5.8688 | 5.3105 | 4.757
| 500 | 51.7 | 26.97 | 18.22 | 13.9 | 11.23 | 9.47 | 8.16 | 7.116 | 6.42 | 5.676
| 600 | 61.4368 | 31.6571 | 21.258 | 16.164 | 13.0362 | 10.976 | 9.442 | 8.2256 | 7.4062 | 6.5886 | Compliance Coefficients
| 700 | 71.0397 | 36.3068 | 24.296 | 18.428 | 14.8182 | 12.482 | 10.724 | 9.3352 | 8.3924 | 7.4336
| 800 | 80.4284 | 40.911 | 27.334 | 20.692 | 16.5966 | 13.988 | 12.006 | 10.4448 | 9.3786 | 8.2686 | CC[0] | 6.6892e-05
| 900 | 89.8796 | 45.7497 | 30.372 | 22.956 | 18.4442 | 15.494 | 13.288 | 11.5544 | 10.3648 | 9.1641 | CC[1] | -0.010638
| 1000 | 99.42 | 50.37 | 33.41 | 25.22 | 20.29 | 17 | 14.57 | 12.664 | 11.351 | 10.037 | CC[2] | 0.58093
| 1100 | 109.536 | 55.03 | 36.488 | 27.528 | 22.118 | 18.54 | 15.89 | 13.8052 | 12.3628 | 10.9596 | CC[3] | 19.9319
| 1200 | 119.652 | 59.69 | 39.566 | 29.836 | 23.946 | 20.08 | 17.21 | 14.9464 | 13.3746 | 11.8822 | R2 | 0.99999
| 1300 | 129.768 | 64.35 | 42.644 | 32.144 | 25.774 | 21.62 | 18.53 | 16.0876 | 14.3864 | 12.8048
| 1400 | 139.884 | 69.01 | 45.722 | 34.452 | 27.602 | 23.16 | 19.85 | 17.2288 | 15.3982 | 13.7274
| 1500 | 150 | 73.67 | 48.8 | 36.76 | 29.43 | 24.7 | 21.17 | 18.37 | 16.41 | 14.65
| 1600 | 160 | 78.382 | 51.926 | 39.112 | 31.35 | 26.318 | 22.536 | 19.598 | 17.496 | 15.758
| 1700 | 170 | 83.094 | 55.052 | 41.464 | 33.27 | 27.936 | 23.902 | 20.826 | 18.582 | 16.866
| 1800 | 180 | 87.806 | 58.178 | 43.816 | 35.19 | 29.554 | 25.268 | 22.054 | 19.668 | 17.974
| 1900 | 190 | 92.518 | 61.304 | 46.168 | 37.11 | 31.172 | 26.634 | 23.282 | 20.754 | 19.082
| 2000 | 200 | 97.23 | 64.43 | 48.52 | 39.03 | 32.79 | 28 | 24.51 | 21.84 | 20.19
Compliance Coefficients
[12 empty] | CC[0] | 0.0002318
[12 empty] | CC[1] | -0.027868
[12 empty] | CC[2] | 1.1475
[12 empty] | CC[3] | 24.7456
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.00059764
[12 empty] | CC[1] | -0.058802
[12 empty] | CC[2] | 1.9591
[12 empty] | CC[3] | 28.9147
[12 empty] | R2 | 0.99998
Compliance Coefficients
[12 empty] | CC[0] | 0.0011673
[12 empty] | CC[1] | -0.097961
[12 empty] | CC[2] | 2.774
[12 empty] | CC[3] | 34.2899
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.0023537
[12 empty] | CC[1] | -0.16603
[12 empty] | CC[2] | 3.979
[12 empty] | CC[3] | 38.6319
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.0037001
[12 empty] | CC[1] | -0.23952
[12 empty] | CC[2] | 5.1655
[12 empty] | CC[3] | 44.4692
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.0062798
[12 empty] | CC[1] | -0.36067
[12 empty] | CC[2] | 6.9448
[12 empty] | CC[3] | 46.6124
[12 empty] | R2 | 0.99999
Compliance Coefficients
[12 empty] | CC[0] | 0.010562
[12 empty] | CC[1] | -0.56972
[12 empty] | CC[2] | 9.6535
[12 empty] | CC[3] | 49.4503
[12 empty] | R2 | 0.99999
[13 empty] | 0.010562 | -0.56972 | 9.6535 | 49.4503 | 0
[13 empty] | 0.00091614 | 0.035725 | 0.43618 | 1.661 | #N/A
[13 empty] | 0.99999 | 4.3911 | #N/A | #N/A | #N/A
[13 empty] | 372101.5642 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28699672.2035 | 327.7965 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## DALNormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
DA Left Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## DARNormCalData
Compliance Characteristic Adjustment to Normalized Scale at 1000 mL
DA Right Lung PneuView® 2.0 System
|  |  | Relative tolerance on pressure: | 3% | %
|  |  | Absolute tolerance on pressure: | 0.5 | cmH2O
| Compliance setting | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| Ideal Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MAX | 103 | 51.5 | 34.3333 | 25.75 | 20.6 | 17.1667 | 14.7857 | 13 | 11.6111 | 10.5
| Actual Pip | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10
| Ideal Pip MIN | 97 | 48.5 | 32.3333 | 24.25 | 19.4 | 16.1667 | 13.7857 | 12 | 10.6111 | 9.5
Adjustment factor | bar: 0.01 to cmH2O 0.1
Adult Lung: Adjusted intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | 0.01: Compliance Setting
| Volume | 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Compliance Coefficients
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119 | CC[0] | 2.3781e-07
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106 | CC[1] | -0.00011907
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172 | CC[2] | 0.019629
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113 | CC[3] | 8.9895
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | R2 | 0.99999
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 |  | 2.3781e-07 | -0.00011907 | 0.019629 | 8.9895 | 0
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 |  | 6.8777e-08 | 2.6416e-05 | 0.0031764 | 0.11914 | #N/A
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 |  | 0.99999 | 3.1615 | #N/A | #N/A | #N/A
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 |  | 717860.4767 | 17 | #N/A | #N/A | #N/A
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 |  | 28699830.0864 | 169.9136 | #N/A | #N/A | #N/A
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675 | Compliance Coefficients
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311 | CC[0] | 9.4524e-06
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212 | CC[1] | -0.0021852
[12 empty] | CC[2] | 0.1736
[12 empty] | CC[3] | 15.6144
[12 empty] | R2 | 0.99998
Adult Lung: Normalized intrapulmonary pressure developed for each Injection Volume / Compliance Setting pair
| Injected | (L/cmH2O) 0.01: Compliance Setting
| Volume | (L/cmH2O) 0.01: (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 11.7826 | 6.8942 | 5.0665 | 4.065 | 3.5126 | 3.1418 | 2.8008 | 2.6021 | 2.4055 | 2.3344
| 200 | 21.5507 | 11.8307 | 8.376 | 6.4871 | 5.4286 | 4.7043 | 4.1598 | 3.7714 | 3.4364 | 3.2119
| 300 | 31.2341 | 16.7001 | 11.566 | 8.8415 | 7.2941 | 6.2332 | 5.4524 | 4.9078 | 4.4346 | 4.106
| 400 | 41.3408 | 21.5194 | 14.7219 | 11.1958 | 9.1429 | 7.7453 | 6.712 | 5.9947 | 5.3837 | 4.9172
| 500 | 51.2443 | 26.2718 | 17.8437 | 13.5163 | 11.0084 | 9.207 | 8.0046 | 7.0817 | 6.3819 | 5.8113
| 600 | 61.1309 | 31.0408 | 20.9997 | 15.8028 | 12.8235 | 10.6855 | 9.2476 | 8.1522 | 7.2983 | 6.606 | Compliance Coefficients
| 700 | 70.7635 | 35.7262 | 24.0362 | 18.0725 | 14.6555 | 12.1808 | 10.4906 | 9.2227 | 8.2802 | 7.4338
| 800 | 81.0225 | 40.6627 | 27.1921 | 20.4438 | 16.437 | 13.6425 | 11.7335 | 10.2602 | 9.1965 | 8.3113 | CC[0] | 7.2426e-05
| 900 | 90.8752 | 45.2644 | 30.2798 | 22.7134 | 18.2353 | 15.1714 | 13.0262 | 11.3966 | 10.1129 | 9.1391 | CC[1] | -0.011257
| 1000 | 100 | 50 | 33.3333 | 25 | 20 | 16.6667 | 14.2857 | 12.5 | 11.1111 | 10 | CC[2] | 0.58651
| 1100 | 109.4464 | 54.836 | 36.421 | 27.3882 | 21.8151 | 18.1183 | 15.5817 | 13.587 | 12.0668 | 10.9106 | CC[3] | 20.3633
| 1200 | 119.2653 | 59.6386 | 39.5428 | 29.624 | 23.6807 | 19.6539 | 16.8777 | 14.7398 | 13.0715 | 11.7881 | R2 | 0.99998
| 1300 | 129.3083 | 64.324 | 42.6987 | 31.9953 | 25.5798 | 21.1895 | 18.1737 | 15.8432 | 14.0272 | 12.649
| 1400 | 139.2551 | 69.1265 | 45.8205 | 34.3665 | 27.395 | 22.7419 | 19.536 | 16.9631 | 15.081 | 13.5762
| 1500 | 149.2019 | 73.745 | 48.9765 | 36.6531 | 29.3445 | 24.2944 | 20.7988 | 18.1653 | 16.0694 | 14.4536
| 1600 | 159.1487 | 78.581 | 52.0846 | 39.021 | 31.237 | 25.9039 | 22.1545 | 19.3116 | 17.1494 | 15.4106
| 1700 | 169.0955 | 83.2664 | 55.3463 | 41.5075 | 33.0286 | 27.4126 | 23.5764 | 20.4578 | 18.1967 | 16.3675
| 1800 | 179.0423 | 87.9685 | 58.7445 | 43.96 | 35.0218 | 29.039 | 24.9321 | 21.7029 | 19.2113 | 17.3576
| 1900 | 188.9891 | 92.7209 | 62.0232 | 46.3279 | 36.9647 | 30.6485 | 26.3043 | 22.9315 | 20.2913 | 18.3311
| 2000 | 198.9358 | 97.6908 | 64.8755 | 48.7297 | 38.7563 | 32.2077 | 27.7262 | 25.0165 | 21.4368 | 19.3212
Compliance Coefficients
[12 empty] | CC[0] | 0.0002588
[12 empty] | CC[1] | -0.030293
[12 empty] | CC[2] | 1.185
[12 empty] | CC[3] | 25.3711
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.00081187
[12 empty] | CC[1] | -0.074691
[12 empty] | CC[2] | 2.3011
[12 empty] | CC[3] | 27.4413
[12 empty] | R2 | 0.99997
Compliance Coefficients
[12 empty] | CC[0] | 0.001967
[12 empty] | CC[1] | -0.15047
[12 empty] | CC[2] | 3.8352
[12 empty] | CC[3] | 29.1394
[12 empty] | R2 | 0.99996
Compliance Coefficients
[12 empty] | CC[0] | 0.003779
[12 empty] | CC[1] | -0.2522
[12 empty] | CC[2] | 5.5826
[12 empty] | CC[3] | 30.9745
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.0038442
[12 empty] | CC[1] | -0.28391
[12 empty] | CC[2] | 6.4846
[12 empty] | CC[3] | 36.0004
[12 empty] | R2 | 0.99987
Compliance Coefficients
[12 empty] | CC[0] | 0.012231
[12 empty] | CC[1] | -0.64082
[12 empty] | CC[2] | 11.1365
[12 empty] | CC[3] | 29.0366
[12 empty] | R2 | 0.99995
Compliance Coefficients
[12 empty] | CC[0] | 0.020917
[12 empty] | CC[1] | -0.98322
[12 empty] | CC[2] | 15.2757
[12 empty] | CC[3] | 25.0191
[12 empty] | R2 | 0.99993
[13 empty] | 0.020917 | -0.98322 | 15.2757 | 25.0191 | 0
[13 empty] | 0.0027082 | 0.10151 | 1.1916 | 4.3518 | #N/A
[13 empty] | 0.99993 | 10.7279 | #N/A | #N/A | #N/A
[13 empty] | 62339.7338 | 17 | #N/A | #N/A | #N/A
[13 empty] | 28698043.5161 | 1956.4839 | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A
[13 empty] | #N/A | #N/A | #N/A | #N/A | #N/A

## Validation Data
Validation Data
Adult Lung: Measured Lung Volume
| Injected | Compliance Setting
| Volume | (L/cmH2O)
| (mL) | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.07 | 0.08 | 0.09 | 0.1
0
100
200
300
400
500
600
700
800
900
1000
1100
1200
1300
1400
1500
1600
1700
1800
1900
2000